FSSP Performance Checks
Goodrich Summer 2012 Field Project

Modifications
David Delene: 080108	- Written
Phondie Simelane: 120907 - Chart added, added 'avg_channel'
#### FSSP SN 4053-0783-90 Calibration Checks 08/12/2012 : Channel peak (P- Ch) is one-based. Average channel (Cal. Avg CH) is centered on the channel number without a 0.5 offset applied. The equations y=-0.00006737x4+0.00654087x3-0.21861587x2+3.71178039x–12.24371265 is used to related the Bead Size (x) to the Water Eq. Size (y). The standard channel value (y - Std CH) is calculated using the water equivalent size (x - Water Eq. Size) and the fitting equation y=-0.0001079172*x3+0.0058707377*x2+0.409621791*x–0.638911101 for 20 channels and y=-1.4732339*10-5*x3+0.0009442954*x2+0.65006293*x–0.929623896 for 30 channels.. Water equivalent, average size and standard channel are obtainable using the 'avg_channel' command.
Date / (yyyy/mm/dd)	Start time (sfm)	End time (sfm)	MOD	Num. / Ch.	Observed / Peak Channel	Pre-peak Count	Peak Count	Post-peak Count	Bead Size / (µm)	Water Equiv. Size	Calculated Average Channel	Standard Channel
2012/08/29	85080.4	85081.0	0	20	14	5.27	5.27	5.18	40.0	32.59	13.99	15.21
2012/08/29	85456.5	85461.0	0	20	5	28.00	28.09	17.09	15.0	12.91	4.85	5.40
2012/08/29	85708.5	85711.0	0	20	4	42.16	49.16	37.7	10.0	8.88	3.97	3.39
2012/08/14	54227.5	54232.0	0	20	4	18.82	33.09	11.64	9.18	7.99	3.89	2.95
2012/08/14	54501.0	54503.0	0	20	4	1.90	4.75	3.15	10.00	8.88	4.13	3.39
2012/08/14	54690.8	54693.0	0	20	5	53.82	55.68	7.23	11.58	10.37	4.60	4.12
2012/08/14	54883.0	54883.9	0	20	5	4.56	12.00	10.78	18.00	14.81	5.23	6.37
2012/08/14	55092.5	55094.0	0	20	9	21.33	26.20	6.33	22.81	18.07	8.72	8.04
2012/08/14	55534.8	55537.5	0	20	12	2.40	5.10	3.37	30.00	24.39	12.09	11.28
2012/07/06*	40918.0	40919.0	0	30	3	40	110	103	8.0	6.53	3.25	3.35
2012/07/06*	40498.0	40499.0	0	30	9	122	154	75	17.3	14.37	8.87	8.57
2012/07/06*	40758.0	40759.0	0	30	22	81	122	85	42.3	32.97	22.01	21.00

*Data read from DMT Particle Analysis and Display System screen shots.