FSSP Calibration Checks Summary

North Dakota 2008 Polcast2 Field Project

Modifications:

Chris Kruse:  080528 – Written

#### FSSP SN 277-0676-06 (NCAR RAL) Calibration Checks:  Channel peak (P- Ch) is one-based.  Average channel (Cal. Avg CH) is centered on the channel number without a 0.5 offset applied.  The equations y=-0.00006737x4+0.00654087x3-0.21861587x2+3.71178039x–12.24371265 is used to related the Bead Size (x) to the Water Eq. Size (y).  The standard channel value (y - Std CH) is calculated using the water equivalent size (x - Water Eq. Size) and the fitting equation y=-0.0021*x2+0.5753*x–1.503.

| Date | Start [sfm] | End [sfm] | P- CH | Pre-Peak Count | Peak Count | Post-Peak Count | M o d | Bead Size [um] | Water Eq. Size | Cal. Avg CH | Std CH |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 08/05/27 | 7927.2 | 7928.9 | 4 | 7.4118 | 14.7647 | 11.7059 | 0 | 15 | 12.91 | 4.1267 | 5.5737 |
| 08/05/27 | 8472.2 | 8480.4 | 4 | 2.6220 | 3.7683 | 3.1829 | 0 | 15 | 12.91 | 4.0586 | 5.5737 |
| 08/05/27 | 8614.3 | 8626.2 | 4 | 0.5378 | 0.6471 | 0.4706 | 0 | 15 | 12.91 | 3.9594 | 5.5737 |
| 08/05/27 | 8736.3 | 8737.3 | 4 | 14.100 | 20.8000 | 10.5000 | 0 | 15 | 12.91 | 3.9207 | 5.5737 |
| 08/06/09 | 70412 | 70415 | 5 | 55.4828 | 63.3793 | 16.2759 | 0 | 15 | 12.91 | 4.7098 | 5.5737 |
| 08/06/09 | 70470 | 70472 | 5 | 75.6316 | 113.053 | 48.5789 | 0 | 15 | 12.91 | 4.8859 | 5.5737 |
| 08/06/09 | 70520.8 | 70522.5 | 5 | 38.6250 | 51.5625 | 17.6250 | 0 | 15 | 12.91 | 4.8052 | 5.5737 |
| 08/06/09 | 70549.5 | 70550.5 | 5 | 31.7778 | 44.0000 | 11.4444 | 0 | 15 | 12.91 | 4.7668 | 5.5737 |
| 08/06/09 | 70693 | 70695 | 11 | 22.5789 | 24.4211 | 23.7895 | 0 | 31.5 | 25.87 | 11.017 | 11.972 |

*Unclear peak found during the given time period.