FSSP Calibration Checks Summary
 Saudi Arabia 2007 Rainfall Augmentation Field Project

Modifications:
  David Delene: 080125 – Written
  David Delene: 080205 – Added Chris's results.
  David Delene: 080217 – Added Water Eq & std CH columns.
  David Delene: 080307 – Denoted bead test used for project calibration.

#### FSSP SN 1947-0281-60 (WMI) Calibration Checks: Channel peak (P- Ch) is one-based. Average channel (Cal. Avg CH) is centered on the channel number without a 0.5 offset applied. The equations y=-0.00006737x4+0.00654087x3-0.21861587x2+3.71178039x–12.24371265 is used to related the Bead Size (x) to the Water Eq. Size (y). The standard channel value (y - Std CH) is calculated using the water equivalent size (x - Water Eq. Size) and the fitting equation y=-0.0021*x2+0.5753*x–1.503.
Date	Start [sfm]	End [sfm]	P- CH	Pre-Peak Count	Peak Count	Post-Peak Count	M / o / d	Bead / Size [um]	Water / Eq. / Size	Cal. / Avg / CH	Std / CH
07/12\/05	46700.0	46703.0	6	10.700	18.233	15.700	0	20.0	16.09	6.112	7.210
07/12/05	49355.5	49358.5	4	8.933	19.800	19.667	0	15.0	12.91	4.212	5.574
08/01/07	27869.7	27870.7	4	35.900	67.300	58.600	0	14.5	12.57	4.140	5.400
08/01/07	28062.0	28063.5	4	18.200	28.530	17.800	0	14.5	12.57	3.994	5.400
08/01/07	28446.2	28447.2	4	39.100	63.300	62.000	0	15.0	12.93	4.139	5.574
08/01/07	28489.0	28490.0	4	3.900	4.500	2.700	0	15.0	12.93	3.900	5.574
08/01/07	28935.0	28936.0	9	52.700	58.200	41.700	0	30.0	24.39	9.928	11.280
08/01/29	31227.7	31228.7	4	22.400	37.500	29.100	0	15.0	12.93	4.075	5.574
08/01/29	31373.5	31376.0	4	28.280	35.160	17.720	0	15.0	12.93	3.870	5.574
08/01/29*	31487.5	31488.5	9	4.100	5.200	4.500	0	30.0	24.39	9.029	11.280
08/02/06	25548.5	25549.5	4	14.100	16.100	10.000	0	15.0	12.93	3.898	5.574
08/02/06	25736.7	25737.7	8	13.300	15.600	12.200	0	30.0	24.39	7.973	11.280
08/02/06	29680.6	29682.0	15	26.571	36.857	19.571	3	8.0	6.53	14.916	16.874
08/03/10	39292.8	39294.9	4	64.191	79.619	10.523	0	15.0	12.91	3.562	5.574
08/03/10	40086.8	40087.5	5*	44.857	64.571	43.000	0	15.0	12.91	4.988	5.574
08/03/10	40257.7	40258.8	7*	45.364	55.000	44.091	0	30	24.39	6.991	11.279

*Unclear peak found during the given time period.
#### FSSP SN 16702-0789-126 (NRL) Calibration Checks: Channel peak (P- Ch) is one-based. Average channel (Cal. Avg CH) is centered on the channel number without a 0.5 offset applied. The equations y=-0.00006737x4+0.00654087x3-0.21861587x2+3.71178039x–12.24371265 is used to related the Bead Size (x) to the water Eq. Size (y). The standard channel value (y - Std CH) is calculated using the water equivalent size (x - Water Eq. Size) and the fitting equation y=-0.0021*x2+0.5753*x–1.503.
Date	Start [sfm]	End [sfm]	P- CH	Pre-Peak Count	Peak Count	Post-Peak Count	M / o / d	Bead / Size [um]	Water / Eq. / Size	Cal. / Avg / CH	Std / CH
07/12/05	46919.5	46920.5	17	24.600	36.100	33.600	0	20.0	16.09	17.095	7.210
07/12/05	47119.0	47122.0	5	53.067	97.667	72.200	0	8.0	6.53	5.0858	2.164
08/02/06	26906.0	26908.0	7	19.800	22.050	19.400	0	15.0	12.70	6.9935	5.574
08/02/06	27247.2	27249.2	16	3.000	4.300	3.000	0	30.0	24.39	16.000	11.280
08/02/06	27485.0	27486.0	5	42.900	86.000	36.600	0	8.0	6.53	4.962	2.164
08/02/06	27598.0	27599.0	5	8.200	16.800	9.000	0	8.0	6.53	5.024	2.164
08/02/06	28911.5	28915.0	19	5.143	6.286	5.143	3	8.0	6.53	19.000	16.874