GCPEx Instrument Care Checklist

# After every flight

## Clean Optical Probes windows or mirrors

### 2DC

#### Clean mirrors in probe tip arms with a Q-tip wetted with isopropyl alcohol.

### CIP

#### Clean mirrors in probe tip arms with a Q-tip wetted with isopropyl alcohol.

### CDP

#### Clean optics windows with a Q-tip wetted with white vinegar to remove water spots.

#### Use isopropyl alcohol or acetone to finish cleaning.

### HVPS-3

#### Visually inspect windows for particulate contamination with LED light.

#### Blow off any particulate contamination with compressed air, taking care not to spray refrigerant onto windows.

#### Lightly swirl optical quality cleaning tissue folded over a Q-tip wetted with anhydrous methanol over windows.  Do not rub.

### DMT Backscatter Probe (BCP)

#### Clean window with a Q-tip wetted with white vinegar to remove water spots.

#### Use isopropyl alcohol or acetone to finish cleaning.

## Blow out Gust Probe Pitot lines

## Purge HVPS-3 can with Nitrogen

## Fill CPC with Butanol

# Once a week

# When needed

## Remove Edgetech sensor head and clean mirror.